| Pasteura 3, 02-093 Warsaw, Poland, tel: (48-22) 589 2213; fax: (48-22) 822 53 42 |  |
| --- | --- |

_____________________________________________________________________________Warsaw, 24.10.2017

Request for proposals no 86/2017

for the supply of

CMOS camera with computer

to

NENCKI INSTITUTE OF EXPERIMENTAL BIOLOGY,

of the Polish Academy of Sciences

Procuring Party: Nencki Institute of Experimental Biology,

PAS with its registered office at ul. Pasteura 3, Warsaw (02-093), Tax ID No (NIP):525-000-92-69, Statistical ID No (REGON): 000325825

Contact person for matters related to supply: Piotr Michaluk

e-mail address: p.michaluk@nencki.gov.pl

- Description of the Subject of the Procurement:

The subject of the procurement is: CMOS camera with computer for image acquisition.

- Description of Offer Preparation, and its Assessment:

The Offer must be drawn up on template 1 enclosed to the Request. The Offer must contain the following:

- Information on total gross and net value of the procurement: The price offer should include: camera with sCMOS sensor, with resolution of 2048 x 2048 pixels and pixel size 6.5 m x 6.5 m. Camera must allow for fast acquisition of 100 frames per second (fps) at full frame size. Camera must have C-mount enabling attachment to a microscope and a possibility of external triggering. Camera sensor should have at least 30 000:1 dynamic range and full-well capacity should be at least 30 000 e-. The kit needs to have PCI card for enabling data transfer from camera to a computer at rate 100 fps. The computer should have 32 GB (or more) of RAM, have Windows 8 (64 bit), or later, operation system and should have all drivers for the camera. It should be possible to operate the camera with Micro-Manager software.

- Deadline for the execution of the procurement

- Proposals must be sent electronically as a photocopy of the original proposal via e-mail at:  p.michaluk@nencki.gov.pl

- Please mark the proposal in the subject of the message: CMOS camera and PC

- Only the proposals containing all points listed above will be considered.

- If the procurement meets the technical requirements listed in Detailed description, the Procuring Party considers the price criterion while making the choice.

- Additional information:

- An agreement will be signed with the selected Contractor for the purpose of the execution of the procurement.

- The Procuring Party reserves the right to negotiate the agreement with the best Bidders.

- The Procuring Party reserves the right to choose none of the Bidders.

Deadline for submitting proposals: not later than 29.10.2017 r., at 15:00

Appendix 1: Form Template

Name and address of the Contractor………………………………………………………………………………….

| No | Deatiled DESCRIPTION | Requirement met (YES/NO) | PRICE | Comments |
| --- | --- | --- | --- | --- |
| 1. | Camera with sCMOS sensor, 2048x2048 pixel resolution, enabling fast acquisition with rate 100 fps at full frame |  |  |  |
| 2. | Pixel size: 6,5 m x 6,5 m |  |  |  |
| 3. | C-mount |  |  |  |
| 4. | Posibility of external triggering |  |  |  |
| 5. | Dynamic range: 30 000:1 or greater |  |  |  |
| 6. | Full-Well capacity: 30 000e- or greater |  |  |  |
| 7. | PCI card for enabling data transfer from camera to a computer at rate 100 fps |  |  |  |
| 8. | The computer should have 32 GB (or more) of RAM, have Windows 8 (64 bit), or later, operation system and should have all drivers for the camera. |  |  |  |
| 9. | Possibility of camera operation using Micro-Manager software |  |  |  |
| 10. | 12-month warranty or longer |  |  |  |
| 11. | Delivery: no longer than 4 weeks |  |  |  |

Procurement execution deadline:……………………………………….

Warranty period:…………………………………………………..…………….

……………………………………………                                …….……………………………………

Place, date                                                 Signature and seal of the Contractor