**Appendix No. 1: Bid form template**

Name and address of the Contractor:……………………………………………………

Contact person..…………………………………Phone number………………………………… e-mail: ………………………………………………….

Subject of the contract: **Equipment of a fluorescence microscope - micromanipulators.**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | **Description of requirements** | **Requirement met (YES/NO)**  | **Device model and manufacturer**  | **Comments** |
| 1. | **Equipment of a fluorescence microscope - micromanipulators.*** Two micromanipulators with X, Y and Z motors suitable for patch-clamp experiments
* Drift (**Long Term Stability)** < 1.0 micron in 2 hours
* < 30 nm resolution
* Electrically quiet (noise free)
* Option of an „approach” mode either by combination of movement of 2 motors - virtual 4th axis, or by placement of a motor along the pipette axis
* Possibility of memorizing pipette positions, for example “home” and “work”
* Travel ≥ 20mm in each axis
* Maximum travel speed ≥ 4 mm/sec
* Controlling cube with a small footprint and a switch to control two manipulators
* 19” rack- mountable control unit which can control 2 micromanipulators
* Mounting posts which enable manipulators to be mounted above the translational platform or brackets, so pipette can reach microscope chamber which is mounted 197 mm above the surface of the platform
 | (YES/NO) (YES/NO)  |  |  |

Net PLN …………………………………………… price,

Shipping ……………………….. net PLN,

The bid price must be expressed in Polish zloty (PLN).

The settlement with the contractor will be made in Polish zloty (PLN).

Contract completion date .………….. from the date of signing the agreement (max. 5 weeks).

Warranty:min. ……………………. (min. 24 months).

I declare that I am not subject to exclusion from the proceedings pursuant to Article 7 paragraph 1 of the Act of 13 April 2022 on special solutions to prevent the support of aggression against Ukraine and to protect national security (Journal of Laws of 2021, item 835)

I have read the model agreement attached as Appendix No. 2 to the request for quotation and I do not raise any objections to it.

 …………………………………………… …………………......………………………… Place, date Signature and stamp of the Contractor

Pasteur 3, 02-093 Warsaw, e-mail: p.michaluk@nencki.edu.pl; <http://www.nencki.edu.pl_>