Warsaw, 02.03.2020

Request for quotation No 27/2020

for the delivery

A set for non-invasive recording of cell electrical activity and analysis of data from multi-electrode matrices

to the Marceli Nencki Institute of Experimental Biology

Polish Academy of Sciences

Ordering party: Institute of Experimental Biology named after M. Nencki, Polish Academy of Sciences, with its registered office at 3 Pasteura Street, Warsaw (02-093), NIP: 525-000-92-69, REGON: 000325825

Contact person for the contract: Katarzyna Kalita-Bykowska

e-mail: k.kalita@nencki.edu.pl

Deadline for submission of tenders: no later than 17.03.2020, to 12:00 p.m.

I. Description of the subject matter of the contract:

The subject of the contract is:

A set for non-invasive recording of cell electrical activity and analysis of data from multi-electrode matrices

- A workstation for recording and analyzing data from one 60 electrode microarray with features:

- Recording and analysis of data from at least 60-electrode microarray and at least 8 additional analogue channels

- There should be a stimulator in the workstation, allowing selective stimulation of selected electrodes through the built-in software

- The set should include an integrated amplifier with a signal filtering system and at least a 2-channel stimulus pulse generator

- Integrated workstation heating system

- The station is capable of operating in an incubator with controlled levels of oxygen and carbon dioxide

- Possibility of electrical recording/stimulation with simultaneous imaging under the inverse microscope in transient light and wide field fluorescence

- Possibility to extend the set with additional workstations used in in vitro /ex vivo experiments

- Possibility of expanding the system with similar units (workstations)

- Option to extend the system with a perfusion system

- The set should include software for data recording and analysis with the possibility of free updates

- Data rate (USB) not less than 50 kHz/channel

- Free technical assistance for at least 2 years

- Multi-electrode microarrays compatible with a reusable workstation and a simple and quick method of sterilisation (including autoclaving)

- Features of multi-electrode microarrays:

- the option of installing dedicated membranes, permeable to oxygen and carbon dioxide and not permeable to water - reducing the risk of liquid evaporation and biological infection during experiments

- possibility of simultaneous high-resolution imaging in inverted microscopy and electrical stimulation

- Transparent bottom for high-resolution inverse microscope imaging

- Analysis of nerve cell electrical activity documented by publications available in PubMed database (at least 30 publications from 2015)

- Life time electronic contact with experts with laboratory experience in the use of equipment

- Warranty: min. 2 years

- Delivery, installation and training

II Criteria for the evaluation of tenders

- Price- 60%.  - (max - 60 points.  Points are awarded according to the formula: (price of the most advantageous offer / price of the examined offer) x 60 points. In the case of bids of foreign Economic Operators, in order to compare bids - the Contracting Authority shall add the applicable VAT to the net price indicated in the offer)

- Publicly documented results obtained from 2015 using the set described in the documentation - Number of publications – 20%  (30  publications- 0 points,  31-100 publications- 10 points, over 100 publications – 20 points )

- Warranty period– 20%: (0 points – 2 years, 10 points - 3 years, 20 points – 4 years)

III Description of the Preparation of the Offer and its Evaluation:

- The offer should be prepared on model No. 1 attached to the Inquiry.

- The offer should include information about the total net and gross value of the order: Before signing the contract, the Economic Operator, whose bid will be selected, will deliver the following scans: certificates of entry into the business register, REGON certificate and NIP certificate.

- Tenders must be sent electronically in the form of a scanned original tender by e-mail to the following address: k.kalita@nencki.edu.pl

Please mark the offer in the message title: A set for non-invasive recording of cell electrical activity and analysis of data from multi-electrode matrices

- Only offers which meet all requirements specified in point I - "Description of the subject matter of the Contract" shall be taken into consideration.

IV Additional information:

- In order to complete the order, a contract will be signed with the selected Contractor.

- The maximum completion date for each order under the contract is max. 9 weeks (the declared delivery date is indicated by the Contractor in the offer).

- The Contracting Authority reserves the right to negotiate the terms of the contract with the best contractors.

- The Contracting Authority reserves the right not to select any of the Economic Operators.

- The choice of the Contractor will be announced on the website. The contracting authority shall do so without delay after the end of the procedure.